## Slide 1: Portable Remote Imaging Spectrometer (PRISM)
P. Mouroulis, B. Van Gorp, R. O. Green
Jet Propulsion Laboratory
California Institute of Technology

## Slide 2
The PRISM Team
Pantazis (Zakos) Mouroulis: PI, optics
Byron Van Gorp: System engineer & mechanics
Robert O. Green: Calibration, radiative transfer
Daniel Wilson: Diffraction grating
Heidi Dierssen (UConn): Science lead
Bo-Cai Gao (NRL): Atmospheric correction algorithm
Joe Boardman (AIG): Orthorectification, geolocation
K. Balasubramanian, David Cohen, Michael Eastwood, Brian Franklin, Linley Kroll, Scott Leland, Frank Loya, Alan Mazer, Ian McCubbin, Doug Moore, Scott Nolte, Otto Polanco, David Randall, Brandon Richardson, Jose Rodriguez, Chuck Sarture, Eugenio Urquiza, Rudy Vargas, Victor White, Karl Yee
Instrument team
Science team

## Slide 3
spectrometer
vacuum enclosure
baseplate
telescope
2-channel SWIR radiometer
PRISM schematic

## Slide 4
Spectral | Range | 350-1050 nm
 | Sampling | 2.85 nm
 | Resolution (FWHM) | 3.5 nm
 | Calibration uncertainty | < 0.1nm
Spatial | Field of view | 31.7o
 | Instantaneous FOV sampling | 0.882 mrad
 | IFOV resolution (FWHM) | 0.97 mrad
 | Cross-track spatial pixels | 608
 | Ground resolution | 0.35 – 20 m
Radiometric | Range | 0 to max. beach R
 | Sampling | 14 bit
 | Stability | >99%
 | Calibration uncertainty | <2%
 | SNR | 2000 @450 nm*
 | Polarization variation: | < 1%
Uniformity | Spectral cross-track uniformity | >95%
 | Spectral IFOV mixing uniformity | >93%
* At AVIRIS equivalent integration time and spectral sampling Values in red are measured
Spectrometer parameters

## Slide 5
Parameter | Channel 1 | Channel 2
Channel center (nm) | 1240* | 1610*
Bandwidth (nm) | 20* | 60*
FOV (mrad, FWHM) | 2.4 | 2.4
Boresight knowledge (mrad, relative to spectr.) | 0.05 | 0.05
Radiometric stability | 99% | 99%
SNR (@4.3 mW/cm2 sr) | 300* | 100*
SWIR channel parameters
* Tentative values, SWIR characterization not completed.

## Slide 6
Performance metric | System parameter | Enabling technology
SNR | Throughput (F/No=1.8) | Dyson spectrometer and two-mirror telescope design
Dynamic range | Readout rate (176 Hz) | Teledyne HyViSI detector and JPL readout electronics
Polarization insensitivity | Angles of incidence, grating response | Optical design, grating groove design, coatings
Uniformity | Optical prescription | Dyson design, concave grating, lithographic slit
Stray light suppression | Optical surfaces, slit, grating | Coating and filter design, black Si slit, low scatter grating
Calibration stability | Operating temperature (~22o C), vibration response | Optomechanical design, detector mount, athermalization
PRISM integrates new technologies and design techniques to achieve performance

## Slide 7
PRISM spectral characteristics: spectral response functions
350-450 nm
650-750 nm
950-1050 nm

## Slide 8
Spectrum of Er-doped spectralon panel received through PRISM (blue curve) and through reference ASD spectrometer (red curve).
Test spectrum extraction

## Slide 9
Scatter plot of spectral channel centroids as a function of spatial location for three isolated wavelengths (Hg lamp 437 nm and 547 nm, and laser at 639 nm).
Spectral uniformity

## Slide 10
Polarization variation throughout the spectral range for five positions spanning the field of view.
Variation = (Imax– Imin)/(Imax+ Imin)*100
Polarization characteristics

## Slide 11
Along-track spatial response functions (without flight motion blur) for all field positions.
Example along-track spatial response functions (without flight motion blur) for four wavelengths spanning the PRISM spectral range at a single field position.
FWHM =~1.1 pixel
Spatial characteristics: Along-track response

## Slide 12
Example cross-track spatial response functions for four wavelengths spanning the PRISM spectral range at a single field position.
FWHM resolution =~1.05 pixel
Spatial characteristics: Cross-track response
Spatial/spectral IFOV mixing

## Slide 13
Radiometric characteristics: Linearity
R2 values for linear fit 0-80% of range
Log scale
Linear scale

## Slide 14
Predicted in-flight PRISM SNR incorporating measured data from known radiance of NIST-traceable lamp illuminated spectralon panel.
Integrating sphere for spectral and radiometric field calibration
Predicted flight SNR based on lab calibration

## Slide 15
Image of several laser lines covering the entire field of view.
(image taken at room temperature, small residual imperfections disappear at operating temperature)
Stray light assessment
Order-sorting filter seam location

## Slide 16
Example orthorectification with camera model
raw corrected
Slit curvature projected on ground (corrected in post-processing)
Chart: 
| -14.4 | -11.2 | -6 | 0 | 6 | 11.2 | 14.4
 | -0.61635 | -0.35359 | -0.080696 | 0 | -0.16788 | -0.51856 | -0.82886
Agrees well with theoretical prediction of 0.82o

## Slide 17
Boresight alignment of SWIR radiometer (blue) and spectrometer (red)
Cross-track, spectrometer channels 321-323
Along-track, without motion blur
SWIR geometry and spatial characteristics

## Slide 18
Aircraft mounting plate
Vibration isolators
INS/GPS unit
Connector plate
Detector electronics
Heater strips
22”
SENSOR HEAD AND AIRCRAFT MOUNTING PLATE

## Slide 19
Operator rack
Control rack
Internal thermal control
Housekeepingdata logging
External heater PS
External heater control
Data recorder
Computer
Keyboard & monitor
Shutter control
ELECTRONICS

## Slide 20
1.5TB portable solid state drive (3x512 GB)
“Hotel kit” allows downloading speed equal to or slightly faster than recording speed, giving same day data access to investigators (4-lane SATA to RAID array), 4-5x faster than download speed to typical computer hard drive.
RAPID DATA TRANSFER
Remove drive
Insert into download module
Read into portable computer

## Slide 21
TEST FLIGHT SCHEDULE
Calibration/engineering flight May 2012:
  Ivanpah Playa (spatial properties, radiometry high-R targets)
  Lake Tahoe (low R water target)
First test/science flight July 2012:
  Monterey Bay/Elkhorn Slough: seagrass coastal habitat
From left to right, Ivanpah Playa calibration site marked with tarps, dark tile targets, portable solar radiometers.
Lake Tahoe buoy